## Slide 1: Breast biomechanical modelling for the evaluation of the compression and the discomfort perception in mammography.
Industrial advisors:
Dr. Serge Muller
Dr. Ann-Katherine Carton
by Mîra Anna.
5 July 2018
Thesis title:
Academic advisor:
Prof. Yohan Payan

## Slide 2: Clinical Background
X-ray tube
X-ray
beam
Image receptor
Exam steps :
Breast positioning
Progressive compression
Image acquisition
Breast compression
Reduce the patient motion
Reduce the absorbed dose
of ionizing photons
Separate overlapping
structures
reduce image
degrading scatter
Is a source of discomfort and sometimes pain
Notes: The breast is positioned on the breast support which is also the image receptor.
The technologist will flatten the breast using a compression paddle until a nearly uniform breast thickness is reached.
The compression paddle is blocked in order to acquire the image
Women with Hight probability to develop breast cancer will participate to a breast cancer screening program. Each 2 years invited to take a mammography exam

## Slide 3
Goal:
Investigate alternative breast compression methods to
Reduce the pain
Decrease the average glandular dose
Improve image quality

## Slide 4: Technical approach
Source
Detector
Numerical breast phantom
Patient specific breast geometry
Psychological and social factors:
  social status
  physical condition
  anxiety
Physical factors:
  compression force
  strain intensity
Develop a patient specific breast biomechanical model
Image simulation framework
Patient comfort
Average Glandular Dose
Image quality
Notes: A cylidrical geometry was used, to perfor a study comparing the breast comression techniques a more complex breast geometry is needed
Due to physical and social factors the patient comfort is difficult to asses even in a clinical study, the results will depend on pain interpetation of each patient, and pain hystory

## Slide 5: Data acquisition
Breast compression data - from the most recent patient mammography
 | Subject 1 |  | Subject 2 | 
 | Right breast | Left breast | Right breast | Left breast
Thick. (mm) | 47 | 42 | 50 | 49
Force (N) | 21.9 | 40.9 | 97.8 | 56.6
MRI images of 2 volunteers
Volunteer 1
Volunteer 2

## Slide 6: Biomechanical breast modelling
Inframammary ligament
Medial ligament
Boundary conditions
Fixed nodes in Y direction
Fixed nodes in Z direction
Muscle mesh nodes
Fascia mesh nodes
3
2
4
1. Mesh components 3. Dirichlet boundary conditions
2. Breast support matrix 4. Boundary conditions
a) Image registration
b) Image segmentation
c) Geometry extraction
Muscle
Fascia
Breast
Skin
1

## Slide 7
Breast compression modelling
Compression paddle models
Rigid paddle (SRP)
Rigid bodies models.
Frictionless contact.
SRP have only 1 degree of freedom.
SFP have 2 degrees of freedom
  Rotational only joint element.
Compression process
Breast geometry in prone configuration –modelling breast positioning step.
Initial stress (gravity loading) is neglected.
Two steps for breast compression:
Breast support is positioned at breast inframammary level.
Compression paddle is lowered until the target breast thickness is reached .
1.
2.
Flex paddle (SFP)

## Slide 8
Model calibration step
Literature data
Constitutive models for breast compression
Young’s modulus (kPa) | Muscle | Breast | Fascia | Skin
Subject 1 | 10 | 0.3 | 120 | 4
Subject 2 | 10 | 0.5 | 160 | 10
Young’s modulus (kPa) | Muscle | Breast | Fascia | Skin
 | 10 | 0.3 | 120 | 4
 | 10 | 0.5 | 160 | 10
Nearly-incompressible Neo – Hookean models Poisson ratio = 0.499
Breast flattening as function of the applied force.
 | Subject 2
 | 50 mm
F | 97.8 N
 | Subject 2
 | 50 mm
F | 97.8 N
Second volunteer – right breast
Compression force – 10 times lower
[1] Mîra, A., Carton, A. K., Muller, S., & Payan, Y. (2018). Breast Biomechanical Modeling for Compression Optimization in Digital Breast Tomosynthesis. In Computer Methods in Biomechanics and Biomedical Engineering (pp. 29-35). Springer, Cham.
[2] Gefen, A., & Dilmoney, B. (2007). Mechanics of the normal woman's breast. Technology and Health Care, 15(4), 259-271.
[3] Rajagopal, V. (2007). Modelling breast tissue mechanics under gravity loading (Doctoral dissertation, ResearchSpace@ Auckland).
Constitutive models

## Slide 9: Constitutive models for breast compression
In-vivo measured breast flattening curves as function of the applied force.
In-vivo measures
Gent constitutive model
Simulated breast flattening curves as function of the applied force .
[4] de Groot, J. E., Branderhorst, W., Grimbergen, C. A., den Heeten, G. J., & Broeders, M. J. (2015). Towards personalized compression in mammography: a comparison study between pressure-and force-standardization. European journal of radiology, 84(3), 384-391.

## Slide 10
Constitutive models for breast compression
Constitutive models
Young’s modulus (kPa) | Muscle | Breast | Fascia | Skin | Jm
Volunteer 1 | 10 | 0.3 | 120 | 4 | 1
Volunteer 2 | 10 | 0.5 | 160 | 10 | 2
Nearly-incompressible Gent models :
Poisson ratio = 0.499
Model calibration step
Literature data
Breast flattening as function of the applied force.
 | Subject 2
 | 50 mm
F | 97.8 N
 | Subject 2
 | 50 mm
F | 97.8 N
 | Subject 1
 | 47 mm
F | 41.9 N
 | Subject 1
 | 47 mm
F | 41.9 N
[1] Mîra, A., Carton, A. K., Muller, S., & Payan, Y. (2018). Breast Biomechanical Modeling for Compression Optimization in Digital Breast Tomosynthesis. In Computer Methods in Biomechanics and Biomedical Engineering (pp. 29-35). Springer, Cham.
[2] Gefen, A., & Dilmoney, B. (2007). Mechanics of the normal woman's breast. Technology and Health Care, 15(4), 259-271.
[3] Rajagopal, V. (2007). Modelling breast tissue mechanics under gravity loading (Doctoral dissertation, ResearchSpace@ Auckland).

## Slide 11: Breast compression quality metrics
with clinical acquisition parameters.
Two sets of microcalcifications:
200m and 300 m in diameter.
Dance et al. model
  AGD = K g c s
K – incident air kerma
g, c, s conversion factors
Strain/stress internal distribution.
Pressure at skin surface
Average Glandular dose
Patient comfort
SNR: background signal to noise ratio.
SDNR: µ-calc – background signal difference to noise ratio.
Clinical acquisition parameters.
Breast thickness
Homogeneous breast composition:
  20% glandularity.
[5] de Carvalho, P. M. (2014). Low-Dose 3D Quantitative Vascular X-ray Imaging of the Breast (Doctoral dissertation, Université Paris Sud-Paris XI).
[6] Dance, D. R., Skinner, C. L., Young, K. C., Beckett, J. R., & Kotre, C. J. (2000). Additional factors for the estimation of mean glandular breast dose using the UK mammography dosimetry protocol. Physics in medicine & biology, 45(11), 3225.
Image quality

## Slide 12: Breast compression quality
Average glandular dose
 | Force (N) | Rigid paddle |  | Flex paddle | 
 |  | BNT(mm) | AGD(mGy) | BNT(mm) | AGD(mGy)
Vol 1 | 22 | 46 | 1.15 | 44 | 1.12
Vol 2 | 95 | 48 | 1.20 | 46 | 1.15
Image Quality
SNR | Rigid paddle |  | Flex paddle |  | 
 | mean | std dev | mean | std dev | p-value
Vol 1 | 82.90 | 43.09 | 83.70 | 37.72 | 0.706
Vol 2 | 126.89 | 8.75 | 137.21 | 10.73 | 0.000
SDNR | Rigid paddle |  | Flex paddle |  | 
 | mean | std dev | mean | std dev | p-value
Vol1 | 0.74 | 0.68 | 0.79 | 0.54 | 0.689
Vol2 | 1.14 | 0.57 | 1.13 | 0.53 | 0.885
SDNR | Rigid paddle |  | Flex paddle |  | 
 | mean | std dev | mean | std dev | p-value
Vol 1 | 2.01 | 1.28 | 1.85 | 1.02 | 0.224
Vol 2 | 0.76 | 3.15 | 0.92 | 0.093 | 0.885
200m microcalcifications
300m microcalcifications
For the two volunteers the flex paddles preserve or improve image quality compared to rigid paddles.
The compressed breast nominal thickness (BNT) and average glandular dose (AGD) are similar with both paddle types.

## Slide 13: Breast compression quality
Subject 1 : patient comfort
0
0.5e-3
0.1e-2
0.003
0.006
0.01
0.05
0.17
MPa
Stress intensity
Strain intensity
0
0. 2
0. 4
0. 6
0. 7
0. 8
0. 9
1.
Skin pressure at paddle contact
0
0. 01
0. 02
0. 03
0. 04
0. 06
0. 08
MPa
Skin pressure at receptor contact
0
0. 01
0. 02
0. 03
0. 04
0. 06
0. 08
MPa

## Slide 14: Breast compression quality
Subject 2: patient comfort
0
0.2e-3
0.2e-2
0.6e-2
0.01
0.02
0.03
0.5
MPa
Stress intensity
0
0.4
0.6
0.8
1
1.2
1.3
1.4
Strain intensity
0
0. 01
0. 02
0. 03
0. 04
0. 05
0. 06
MPa
Skin pressure at receptor contact
0
0. 01
0. 02
0. 03
0. 04
0. 05
0. 06
MPa
Skin pressure at paddle contact

## Slide 15: Conclusion
For the small breast volume, the impact of the flex paddle on image quality, AGD or patient comfort is non significant.
For the large breast volume, the flex paddle leads to improved image quality and patient comfort, with no significant impact on AGD.

## Slide 16: Outlook
Improve the soft tissue constitutive models .
Introduce breast heterogeneity and viscosity.
Apply the breast compression using a larger number of breast models.
Compare the compression results with a larger set of in-vivo measured data (Mammography or tomography images) .
Develop a new measure of AGD including variable breast thickness.

## Slide 17: thanks!
Any questions?
You can find me at
anna.mira@univ-grenoble-alpes.fr

## Slide 18
Stress-free breast geometry

## Slide 19
Elastic parameters optimization